Supplementary Note on Arduino Firmware Used

in the TRMS Experiments 1 and 2

This supplementary note describes the firmware used in the two configurations of the thermoelectric real-time measurement system (TRMS). The aim is to clarify the function and structure of the Arduino programs so that readers who wish to reproduce or adapt the instrument can understand the main measurement and communication steps.

1) Listing 1: Core firmware used in experiment 1, showing the Arduino program for the NodeMCU ESP8266 that reads two MAX6675 thermocouple channels, calculates the temperature difference across the thermoelectric module, and transmits the temperature variables to the Blynk dashboard for real-time monitoring.

2) Listing 2: Extended firmware used in experiment 2, showing the Arduino program for the WeMos D1 mini ESP8266 that reads two MAX6675 thermocouple channels, communicates with the PZEM-017 through a MAX485/RS485 Modbus link, and transmits temperature, voltage, current, power, and energy data to the Blynk dashboard for real-time monitoring and logging.

Listing 1:

Firmware structure for Experiment 1: temperature-gradient measurement and monitoring

/****************************************************
Listing 1.
Core firmware structure for Experiment 1:
temperature-gradient measurement and monitoring
Controller: NodeMCU ESP8266
****************************************************/

#include <ESP8266WiFi.h>
#include <BlynkSimpleEsp8266.h>// Library V 1.0.1
#include <MAX6675.h>//Library Robb Tillaat v 0.1.1

/* -------- Pin mapping -------- */
// MAX6675 #1 (hot side)
const int clockPin = D5;
const int dataPinHot = D6;
const int csPinHot = D1;

// MAX6675 #2 (cold side)
const int dataPinCold = D7;
const int csPinCold = D2;

/* -------- Sensor initialization -------- */
MAX6675 thermocoupleHot;
MAX6675 thermocoupleCold;

/* -------- Temperature acquisition and processing -------- */
void readTemperature()
{
float T1 = thermocoupleHot.read();
float T2 = thermocoupleCold.read();
float deltaT = T1 - T2;

Blynk.virtualWrite(V0, T1);
Blynk.virtualWrite(V1, T2);
Blynk.virtualWrite(V2, deltaT);
}

void setup()
{
thermocoupleHot.begin(clockPin, csPinHot, dataPinHot);
thermocoupleCold.begin(clockPin, csPinCold, dataPinCold);
}

Description of the Listing 1:

Listing 1 presents the firmware used in experiment 1, in which the TRMS was configured for real-time measurement of the temperature gradient across the thermoelectric module. The program was developed for a NodeMCU ESP8266 controller and uses the libraries ESP8266WiFi.h, BlynkSimpleEsp8266.h, and MAX6675.h. These libraries provide Wi-Fi communication, transmission of measured values to the Blynk mobile dashboard, and digital acquisition of temperature data from the MAX6675 thermocouple interface modules.

The code is organized into several functional sections. First, the pin-mapping section assigns the digital connections for two MAX6675 modules, one connected to the thermocouple on the hot side of the thermoelectric module and the other to the thermocouple on the cold side. Two MAX6675 objects are then declared to represent these sensor channels. The main acquisition routine, readTemperature(), reads the hot-side temperature T1, reads the cold-side temperature T2, calculates the temperature difference ΔT=T1-T2, and sends these three variables to the Blynk dashboard through virtual pins for real-time display. The setup() routine initializes the two MAX6675 interfaces using the specified clock, chip-select, and data pins.

This firmware therefore implements the essential function of experiment 1: real-time monitoring of hot-side temperature, cold-side temperature, and temperature difference. Its structure is intentionally simple so that readers can readily identify the main program elements and adapt the code for similar classroom measurements or modified sensor layouts.

Listing 2:

Firmware structure for Experiment 2: electrical characterization and extended logging

/****************************************************
Listing 2.
Core firmware structure for Experiment 2:
electrical characterization and extended logging
Controller: WeMos D1 mini ESP8266
****************************************************/

#include <ESP8266WiFi.h>
#include <BlynkSimpleEsp8266.h> // Library V 1.0.1
#include <MAX6675.h> //Library Robb Tillaat v 0.1.1
#include <ModbusMaster.h>//Author:: Doc Walker <4-20ma@wvfans.net>  V 2.0.1
#include <SoftwareSerial.h>

/* -------- Pin mapping -------- */
// MAX6675 #1 (hot side)
const uint8_t thermoSCK = D7;
const uint8_t thermoSO = D6;
const uint8_t thermoCS1 = D2;

// MAX6675 #2 (cold side)
const uint8_t thermoCS2 = D1;

// MAX485 / RS485 interface for PZEM-017
const uint8_t MAX485_RO = D5;
const uint8_t MAX485_DI = D3;
const uint8_t MAX485_RE = D0;
const uint8_t MAX485_DE = D4;

/* -------- Sensor initialization -------- */
MAX6675 thermocoupleHot;
MAX6675 thermocoupleCold;

SoftwareSerial pzemSerial(MAX485_RO, MAX485_DI);
ModbusMaster node;

/* -------- RS485 control -------- */
void preTransmission()
{
digitalWrite(MAX485_RE, HIGH);
digitalWrite(MAX485_DE, HIGH);
}

void postTransmission()
{
digitalWrite(MAX485_RE, LOW);
digitalWrite(MAX485_DE, LOW);
}

/* -------- Measurement acquisition and processing -------- */
void readSystem()
{
float T_hot = thermocoupleHot.read();
float T_cold = thermocoupleCold.read();
float deltaT = T_hot - T_cold;

uint8_t result = node.readInputRegisters(0x0000, 6);
if (result == node.ku8MBSuccess)
{
float voltage = node.getResponseBuffer(0) / 100.0;
float current = node.getResponseBuffer(1) / 100.0;
float power = node.getResponseBuffer(3) / 10.0;
float energy = node.getResponseBuffer(5) / 1000.0;

Blynk.virtualWrite(V0, T_hot);
Blynk.virtualWrite(V1, T_cold);
Blynk.virtualWrite(V2, deltaT);
Blynk.virtualWrite(V3, voltage);
Blynk.virtualWrite(V4, current);
Blynk.virtualWrite(V5, power);
Blynk.virtualWrite(V6, energy);
}
}

void setup()
{
pinMode(MAX485_RE, OUTPUT);
pinMode(MAX485_DE, OUTPUT);

thermocoupleHot.begin(thermoSCK, thermoCS1, thermoSO);
thermocoupleCold.begin(thermoSCK, thermoCS2, thermoSO);

pzemSerial.begin(9600);
node.begin(1, pzemSerial);
node.preTransmission(preTransmission);
node.postTransmission(postTransmission);
}

Description of the Listing 2:

Listing 2 presents the firmware used in experiment 2, in which the TRMS was extended from temperature monitoring to simultaneous thermal and electrical characterization. The program was developed for a WeMos D1 mini ESP8266 controller and uses the libraries ESP8266WiFi.h, BlynkSimpleEsp8266.h, MAX6675.h, ModbusMaster.h, and SoftwareSerial.h. Together, these libraries support Wi-Fi communication, mobile dashboard transmission through Blynk, thermocouple measurement, Modbus communication, and serial communication with the PZEM-017 power measurement module through a MAX485-based RS485 interface.

As in listing 1, the code begins with pin mapping for the thermocouple channels, but it additionally defines the communication pins used by the MAX485 interface. Two MAX6675 objects are declared for the hot-side and cold-side temperature measurements, while a SoftwareSerial object and a ModbusMaster object are used to communicate with the PZEM-017. The functions preTransmission() and postTransmission() control the RS485 transceiver direction during Modbus communication.

The main acquisition routine, readSystem(), first reads the hot-side temperature Thot, the cold-side temperature Tcold, and calculates the temperature difference ΔT. It then requests electrical data from the PZEM-017 and extracts the measured voltage, current, power, and cumulative energy. These thermal and electrical variables are transmitted to the Blynk dashboard through virtual pins so that they can be displayed simultaneously in real time. The setup() routine initializes the RS485 control pins, the two MAX6675 channels, the serial communication with the PZEM-017, and the Modbus transmission callbacks.

This firmware therefore provides the complete software structure for experiment 2, integrating temperature sensing, electrical measurement, and smartphone-based real-time visualisation within a single ESP8266-based system. It may serve as a useful reference for readers wishing to adapt the TRMS for similar thermoelectric or energy-conversion investigations.